## Slide 1
Packaging of RMD APDs with External Mesh Electrodes
on U-Penn Amplifiers
Bert Harrop
Princeton University
Mar. 17, 2017

## Slide 2
APD
Sintered Au layer (“bottom” side)
Mesh Screen (“top” side)
Kapton Tape, 0.002 in.
- HV connected to “top” (p-side)
Ground
Output to Scope
Contact between screen and n+ side made by Au epoxy thru hole in Kapton
p+
Al2O3 Substrate
n-side contact on original APD
http://physics.princeton.edu/~mcdonald/LHC/Tsang/APD_8x8mm_mesh_screen_111512k.ppt
Sketch by Thomas Tsang
n+
McMaster 9227T417, 200 mesh, 33.6% open
Al mesh screen
connection
2-mil
Kapton
spacing
~ 128 µm
opening
~ 85 µm
(Kapton transmit > 93% 980 nm light)
RMD 8x8 mm2 APD – Al-mesh-screen scenario (Nov. 15, 2012)

## Slide 3
2017: Magnetron Sputtering of Top/Bottom APD Contacts
Bottom Contact (Mesa)
Top Contact (Dot)
A bare RMD APD die is 10x10 mm2.
The sensitive area is about 8x8 mm2.
The top (p-side) is photosensitive.
A small electrical contact is made to the top side.
The bottom (n-side) has a 7x7 mm2 “mesa,” which serves as a large electrical contact.
These contacts are made by sputtering
Ti/Au layers of 20/200 nm.

## Slide 4
Create 7x7 mm2 opening in Mask
Applying the conductive H22E Epotek Silver Epoxy
to the 10x10 mm2 Contact Pad of the Penn PCB
Apply Thin Film (50 m) Adhesive
Apply Silver Epoxy, Remove Mask

## Slide 5
Die Bonding of the APD
Placing APD onto epoxy
Cure Epoxy in Vacuum Oven
5 Minutes @ 150C

## Slide 6
Wire Bond Top Contact (for –HV)
25-m-Al Wirebond (Wedge)

## Slide 7
Electroplated Ni Mesh
A
B
C
5-m sheet of 100 electroplated Ni meshes (Au flashed).
3 Mesh Designs:
A = 8-m bar, 160-m spacing, 90% open
B = 10-m bar, 10-m spacing, 11% open
C = 10-m bar, 40-m spacing, 64% open
We use the type A meshes now.

## Slide 8
Mount Mesh on the APD
Type A Mesh mounted on 50-m Kapton
(100-nm-Au flashed)
Mesh mounted on APD
First transistor of the preamp

## Slide 9
Wirebond Mesh to Amplifier Input
Mesh Signal Wirebonds
-HV Wirebond to top (p-side) contact
This version of the preamp/APD operates with the bottom side of the APD at ground, its top side at –HV, and the mesh at ground.

## Slide 10
Completed APD / Amplifier