## Slide 1
Gain & Time Jitter Study for Wenteq Amplifier
Changguo Lu
7/23/2013

## Slide 2
Test setup
Test APD, VCSEL and its collimator is in an environmental chamber (EC). They are in the dark.
Internal view of the EC.
APD test box (64mm^2 APD)
Optical fiber
Collimation lens (Thorlabs CFC-11X-B)
EC chamber
LeCroy Pulse generator
Agilent digital scope

## Slide 3
Parameters used in the test
Use the pulse coming from a 64mm^2 mesh APD with VCSEL laser pulse to test the gain and time resolution.
The VCSEL laser diode is activated by LeCroy pulser with the following parameters:
V = 1.05V; width = 0.9ns; Leading edge = 0.25ns; Trailing edge = 0.25ns; Frequency = 50kHz.
APD HV = -1750V.

## Slide 4
Chart: 
Fe-55 full energy spectrum
Fe-55 full energy peak at -50mV
Chart: 
Fe-55 full energy peak at -145mV
The peak spectrum of Fe-55 is quite different for mesh APD and plain APD.

## Slide 5
VCSEL trigger signal V = 1.05V
# of MIPs = 576/50*1640/6000 ~3.1 MIPs
Time resolution = 20.1ps
Chart: 
| 2.5950000000000001E-8 | 2.5960000000000002E-8 | 2.597E-8 | 2.5980000000000002E-8 | 2.599E-8 | 2.6000000000000001E-8 | 2.6009999999999999E-8 | 2.6020000000000001E-8 | 2.6030000000000002E-8 | 2.604E-8 | 2.6050000000000002E-8 | 2.606E-8 | 2.6070000000000001E-8 | 2.6079999999999999E-8 | 2.6090000000000001E-8 | 2.6100000000000002E-8 | 2.611E-8 | 2.6120000000000002E-8 | 2.613E-8 | 2.6140000000000001E-8 | 2.6149999999999999E-8
 | 0.46987 | 3.179 | 16.7815 | 69.117 | 222.1041 | 556.8598 | 1089.3124 | 1662.5568 | 1979.7847 | 1839.4029 | 1333.3778 | 754.1313 | 332.7807 | 114.5742 | 30.7775 | 6.4505 | 1.0548 | 0.13458 | 0.013396 | 0.0010405 | 6.3049e-05
 | 0 | 0 | 3 | 45 | 170 | 525 | 1088 | 1704 | 2038 | 1726 | 1357 | 767 | 370 | 159 | 35 | 10 | 2 | 1 | 0 | 0 | 0
Chart: 

## Slide 6
VCSEL pulse p.e. estimation
Since we already calibrated the laser beam power, we can convert the signal amplitude unit to # of photons in each pulse:
Repetition rate of laser diode is R = 10 MHz, 980nm photon energy is E980 = 2.03E-19 J, MIP equivalent # of photo-electron = 6000, transmittance of mesh/Kapton = Xtransmit, and we assume that Q.E. of APD = 35%:
@Trigger pulse = 1.05V, VCSEL output power P(W) =0.48*10-6 W,
The conversion fact from laser beam power P(W) to # of MIP is:
# of MIPs = P(W)/R(Hz)/E980 (J)*QE*Xtransmit /6000
= 0.48*10-6 *0.35*Xtransmit /(107 *2.03*10-19*6000)
= 13.8Xtransmit
We already use Fe-55 source calibrated VCSEL laser beam, @V=1.05V, the equivalent # of MIPs ~3.1 (see previous slide),
3.1 = 13.8Xtransmit
 Xtransmit = 3.1/13.8 = 22.5%, the transmittance is small, but reasonable.

## Slide 7
Why the mesh/Kapton transmittance so small?
Under microscope see a lot of unknown substance between mesh wires

## Slide 8
Sr-90 source test
Chart: V(V)
Two APDs coincidence trigger, top APD is a sick APD, on left is the spectrum for bottom APD.