## Slide 1: APD time resolution study with laser diode
C. Lu
Princeton University
(10/3/2014)

## Slide 2
Mesh APD test circuitry
Pigtailed laser diode
SLDP-980-10-SM
Laser beam output lenses
The box can host two APDs, but in this test we only installed one APD, which is near the side wall. The laser beam thru a hole in the side wall enters the box, and hits the mesh of the APD.
The laser diode is triggered by HP8131A pulse generator.
The 8x8 mm2 APD was serial no. 431-3-A

## Slide 3
Measured Fe-55 spectrum
Measured spectrum for Fe-55, HV @ -1750V.
Chart: V(V)
Fe55 peak is at -102mV, which corresponds to 5900/3.6=1639 e-h pairs. In 60um silicon layer MIP particle can create ~6000 e-h pairs. We should adjust laser intensity to 6000/1639 = 3.7 times Fe55 peak, ~370mV, to mimic MIP signal.

## Slide 4
Laser diode spectrum with Wenteq 5010 amplifier
Chart: 12 bits data
Chart: 11 bits data
Chart: 10 bits data
We used three different scope resolution settings to take the data: 12bits, 11bits and 10bits. The signal sizes are same as seen in three histograms above. Later we’ll show their big influence on time resolution.
The High Resolution acquisition mode reduces noise and improves voltage resolution by averaging waveform points from the same trigger event. This mode runs in hardware on full sample rate data so it is very fast and can run at very large memory depths; however, it does reduce bandwidth so it is a type of low-pass filter which reduces the waveform bandwidth.

## Slide 5
Time resolution with Wenteq 5008 amplifier
Chart: 
| 2.0500000000000002E-8 | 2.0510000000000003E-8 | 2.0520000000000001E-8 | 2.0530000000000003E-8 | 2.0540000000000001E-8 | 2.0550000000000002E-8 | 2.056E-8 | 2.0570000000000002E-8 | 2.0580000000000003E-8 | 2.0590000000000001E-8 | 2.0600000000000002E-8 | 2.0610000000000001E-8 | 2.0620000000000002E-8 | 2.063E-8 | 2.0640000000000002E-8 | 2.0650000000000003E-8 | 2.0660000000000001E-8 | 2.0670000000000002E-8 | 2.0680000000000001E-8 | 2.0690000000000002E-8
 | 0.0005534 | 0.0055273 | 0.043901 | 0.27729 | 1.3928 | 5.5632 | 17.6707 | 44.635 | 89.658 | 143.2168 | 181.9242 | 183.7715 | 147.6241 | 94.3035 | 47.906 | 19.3528 | 6.2171 | 1.5883 | 0.32266 | 0.052128
 | 1 | 1 | 2 | 3 | 8 | 16 | 27 | 58 | 131 | 154 | 188 | 175 | 125 | 60 | 31 | 16 | 4 | 0 | 0 | 0
Chart: 
| 2.0599999999999999E-8 | 2.0610000000000001E-8 | 2.0619999999999999E-8 | 2.063E-8 | 2.0639999999999998E-8 | 2.065E-8 | 2.0659999999999998E-8 | 2.0669999999999999E-8 | 2.0680000000000001E-8 | 2.0689999999999999E-8 | 2.07E-8 | 2.0709999999999998E-8 | 2.072E-8 | 2.0729999999999998E-8 | 2.0739999999999999E-8 | 2.0750000000000001E-8 | 2.0759999999999999E-8 | 2.077E-8 | 2.0779999999999998E-8
Fitting | 0.0058028 | 0.14822 | 2.133 | 17.2947 | 79.0047 | 203.3365 | 294.8496 | 240.8846 | 110.8768 | 28.7538 | 4.2012 | 0.34584 | 0.01604 | 0.00041913 | 6.1704e-06 | 5.1181e-08 | 2.3918e-10 | 6.2975e-13 | 9.3418e-16
Data | 4 | 1 | 3 | 29 | 82 | 195 | 300 | 242 | 105 | 33 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Chart: 
| 2.0199999999999999E-8 | 2.021E-8 | 2.0219999999999998E-8 | 2.023E-8 | 2.0239999999999998E-8 | 2.0249999999999999E-8 | 2.0259999999999997E-8 | 2.0269999999999999E-8 | 2.028E-8 | 2.0289999999999998E-8 | 2.03E-8 | 2.0309999999999998E-8 | 2.0319999999999999E-8 | 2.0329999999999997E-8 | 2.0339999999999999E-8 | 2.035E-8 | 2.0359999999999998E-8 | 2.037E-8 | 2.0379999999999998E-8 | 2.0389999999999999E-8 | 2.0399999999999997E-8 | 2.0409999999999999E-8 | 2.042E-8 | 2.0429999999999998E-8 | 2.044E-8 | 2.0449999999999998E-8 | 2.0459999999999999E-8 | 2.0469999999999997E-8 | 2.0479999999999999E-8 | 2.049E-8 | 2.0499999999999998E-8 | 2.051E-8 | 2.0519999999999998E-8 | 2.0529999999999999E-8 | 2.0539999999999997E-8 | 2.0549999999999999E-8 | 2.056E-8 | 2.0569999999999998E-8 | 2.058E-8 | 2.0589999999999998E-8 | 2.0599999999999999E-8
fitting | 2.5132e-09 | 2.2919e-08 | 1.8838e-07 | 1.3957e-06 | 9.3204e-06 | 5.6101e-05 | 0.00030437 | 0.0014884 | 0.0065607 | 0.026065 | 0.093342 | 0.30129 | 0.87656 | 2.2987 | 5.4333 | 11.5758 | 22.2294 | 38.4771 | 60.0304 | 84.4179 | 107.0022 | 122.2492 | 125.8907 | 116.852 | 97.7628 | 73.7235 | 50.111 | 30.7011 | 16.954 | 8.4388 | 3.7861 | 1.531 | 0.55806 | 0.18335 | 0.054295 | 0.014492 | 0.0034867 | 0.00075611 | 0.00014779 | 2.6038e-05 | 4.1349e-06
data | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 23 | 37 | 41 | 71 | 68 | 96 | 124 | 118 | 133 | 126 | 45 | 47 | 24 | 12 | 17 | 6 | 2 | 3 | 2 | 2 | 0 | 1 | 0 | 0 | 0 | 0
Scope Mode(bits) | 12 | 11 | 10
Rms noise(mV) | 3.3 | 4.1 | 5.6
Time resolution (ps) | 13.2 | 20.9 | 31.0
Although the time resolution for 10 bits data set is much worse than 12 bits, if our analysis algorithm gets smarter, it has big room for improvement as we’ll demonstrate later (fft low pass filtration).

## Slide 6
Waveforms recorded with three different scope resolution
Chart: 
Chart: 
Chart: 
For each resolution I plotted three waveforms. It is clearly seen that the noise increases with lower scope resolution. If we use constant fraction algorithm to analyze the data, the zero crossing point will certainly be affected by this noise.

## Slide 7
Noise reduction by fft technique
Chart: 
| 2.0100000000000001E-8 | 2.0110000000000003E-8 | 2.0120000000000001E-8 | 2.0130000000000002E-8 | 2.014E-8 | 2.0150000000000002E-8 | 2.016E-8 | 2.0170000000000001E-8 | 2.0180000000000003E-8 | 2.0190000000000001E-8 | 2.0200000000000002E-8 | 2.021E-8 | 2.0220000000000002E-8 | 2.023E-8 | 2.0240000000000001E-8 | 2.0250000000000003E-8 | 2.0260000000000001E-8 | 2.0270000000000002E-8 | 2.028E-8 | 2.0290000000000002E-8 | 2.03E-8
 | 0.13169 | 0.86262 | 4.2854 | 16.146 | 46.1353 | 99.9767 | 164.3089 | 204.7956 | 193.5876 | 138.7815 | 75.4541 | 31.1123 | 9.7292 | 2.3074 | 0.41501 | 0.056611 | 0.0058564 | 0.00045948 | 2.734e-05 | 1.2337e-06 | 4.2223e-08
 | 6 | 7 | 6 | 23 | 48 | 106 | 147 | 218 | 188 | 144 | 72 | 30 | 4 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
During the data taking we can choose higher scope resolution to reduce the noise.
After the data taking we also can use fft low pass filter algorithm to reduce the bandwidth in the waveform data file:
Run the data files thru fft filter, which set the cutoff frequency at 500MHz, create a set of new data file, then we run the constant fraction data analysis for the new data files.
As a demonstration we analyzed WENTEQ 5008 10 bits data set, with cutoff set at 500MHz, the time resolution is 19ps (w/o fft it was 31ps)!
Why 500MHz?
The relationship between risetime and bandwidth is given by:
trise = 0.35/ f3dB
If we want to preserve 700ps rise time,
f3dB = 0.35/ trise
= 500MHz

## Slide 8
Input impedance of WENTEQ amplifier
We use the left setup trying to measure the input impedance of the amplifier. Change the value of R, record the pulse amplitude V:
R(ohm) | V(v) | Rx(ohm)
440 | 0.606 | 467
0 | 1.188 | 87.6
1094 | 0.284 | 276
Use each pair of the tests we can derive Rx, but all three values are quite different?! This model is not right, but it is certain that Rx is not 50ohm.
In manufacturer’s specs there is no mention of input impedance for the RF amplifiers so far we have tested.

## Slide 9
Summary of laser diode test time resolution with various amplifiers
Amplifier | Time resolution (ps) (12 bits mode) | Time resolution(ps) (10 bits mode) | 
 |  | w/o fft filter | w/ 500MHz fft
ZX60 | 16.1 | 50 | 23.9
MITEQ | 22.7 | 34.5 | 25.3
WENTEQ 5008 | 13.2 | 31.0 | 19.0
WENTEQ 5010 | 18.4 |  | 

## Slide 10
Summary of the test
For APD application the suitable bandwidth of amplifier is around 500MHz, which can limit high frequency noise as well as preserve most essential fast leading edge;
If MIP signal from APD is ~4000 e/h and 980nm laser beam created uniform distributed photoelectrons in the APD, the time resolution for MIP particle would be ~16ps.