## Slide 1
Ultra-Fast Silicon Detectors
First measurements on 75 m LGAD
Contributions to UFSD time resolution
R&D for ATLAS HGTD by CSM
Hartmut F.-W. Sadrozinski
SCIPP, Univ. of California Santa Cruz, CA 95064, USA
22 Feb, 2016
for the
UFSD Group
UCSC – Torino – CNM Barcelona – IJS Ljubljana- LPNHE

## Slide 2
LGAD Testing with  source
For the first beam tests we used commercial amplifiers with bias-T.
For the next tests with 90Sr sources and beams, we developed a broadband amplifier
(BW > 2 GHz, discrete SiGe frontend, analog readout into digital scope)
Signals in:
LGAD, power, sensor bias
Signals out:
single-ended analog signals into 50 Ω
Important: development of shielding
Use  source as test bed for signal and noise performance:
tune the details of the circuit with LGAD data!.
Precision measurements will be done in beam tests.
Scintillator + SiPM (Trigger)
LGAD Signal

## Slide 3
LGAD detectors for HGTD Beam Test
Expect by May 2016:
2x2 arrays of 3-mm and 2-mm side length
various thickness and doping concentration (see later slides)
Amplification and time digitization is required locally for trigger
ASIC development just starting
For the first beam tests plan to use
broadband amplifier developed by UCSC
4 amplifier channels (~ 1-mm width)
wire-bonded to 2x2 array
2 mm

## Slide 4
 Beam data with 75 m LGAD (G = ?)
Pulse shape suggests gain G ≈ 10!
Time Jitter: 11 ps
Time Jitter: ~ 200 ps
75 m,
‘s
Pulses
shorter
than
simulation!?
Ave. Time
Jitter < 40 ps
300 m,
CERN BT
Filtered time
resolution ~100 ps

## Slide 5
dV/dt Comparison of 75 m – 300 m LGAD
300 m:
21 average used to reduce noise
1 ns shaping ≈ 1/3 of rise time
75 m:
21 averages reduces dV/dt by ½
500 ps shaping ≈ rise time
Fast, low-noise amplifier a requirement
High frequencies not important for 300 m
(Need to adjust the CFD threshold)
High frequencies very important for 75 m
(only for BB)
Amplitude
dV/dt
Amplitude
dV/dt

## Slide 6
C-V 50 m (W8) & 75 m (W11) epi LGAD Run 6827
75 m: C-V suggests gain G > 10!
50 m: C-V suggests gain G < 3!
also measured
300 m:
measured gain G > 10!

## Slide 7
Toward LGAD Specifications for HGTD
Larger cell size, higher gain, lower resistivity (?) ==> increased shot noise,
Larger cell size ==> increased electronic noise (CSA), lower dV/dt (BB)
Thicker sensors ==> increased Landau Noise
Lower input impedance ==> lower trans-impedance
Electronics and sensors are closely coupled.
Both LGAD and no-gain n-in-p are being considered
To get to 10 ps jitter:
dV/dt at threshold > 0.5 V/ps ?
Noise (into 50 Ω) < 5 V ?
LGAD details (all numbers TBD):
Gain 10 - 20
Thickness <100 m
Resistivity 0.1 - 10 kΩ-cm
Cell size dimension 2-5 mm
Impedance at amplifier input 20 – 50 Ω
Radiation Hardness (Leakage current, Noise, dV/dt)
Power per cell
Power per cm2

## Slide 8
Time Jitter = Noise/slope = Noise / (dV/dt)
Hard to predict the noise in the moment.
At least we know that for BB it is constant vs. capacitance.
Need to know what it is as a function of the input impedance.
(Data taken with amplifier with gain = 130 shows 4 V at 1 GHz.)
Jitter

## Slide 9
Assume that the resolution scales with slope dV/dt.
For 50 m LGAD (G=10), required dV/dt = 0.7 V/ps (correcting for gain = 100 in WF2).
Slope dV/dt = 0.5 V/ps would give 40 ps
Slew-rate (Slope) dV/dt

## Slide 10
Bulk Resistivity
  Resistivity 0.1 - 10 kΩ-cm
Low-resistivity bulk (100 Ω-cm) generates high field at the junction
(about ½ what is needed for LGAD).
Only half needs to be supplied by p+ multiplication layer
Observe remarkable bias reach in low-resistivity LGAD
We are operating
50 m & 70 m epi sensors
with 100 Ω-cm bulk
at 650 V bias.

## Slide 11
E-field in LGAD bulk adds to the “gain field”
Low-resistivity bulk (100 Ω-cm) generates high field at the junction
(about ½ what is needed for LGAD).
Only 1/2 is supplied by p+ multiplication layer
Near-term: thin 4” wafers at CNM
SOI 100 m & 150 m > 5 kΩ-cm
Epi 50 m 100 Ω-cm; 75m, 100 Ω-cm; 110 m, 600 Ω-cm

## Slide 12
Landau Fluctuations
Beam Tests show that we can correct for time walk.
So assuming we can beat the time jitter with lower noise and higher gain, Landau fluctuations become the time resolution floor which we can’t go below.
They depend on the sensors thickness, independent of the gain.
Both thin sensors, and low threshold are required for good timing resolution:
Another reason to develop low- noise amplifiers

## Slide 13
  Cell size dimension 2-5 mm
Need gain, need small detector sizes to get to sufficient dV/dt Into 50 Ω
Cell Size

## Slide 14
Lowering the amplifier input impedance allows higher slope dV/dt.
We need to ascertain that the noise scales exactly, too.
Impedance at Amplifier Input 20 – 50 Ω

## Slide 15
WF2 Simulation: Radiation Damage (Trapping)
Weightfield 2 (WF2) Implementation and Simulation: Bianca Baldassari
300 µm,
Gain = 10
50µm,
Gain=10
The rising edge of thin sensors is essentially insensitive to trapping: good for timing!
Effect of trapping in thick and thin detectors at large fluences is different since the trapping distance is ~ 50 µm.
Trapping:
Degradation of pulse height in 4*1015 neq/cm2
down to ~ 20% for 300 µm LGAD
down to ~ 50% for 50 µm LGAD
An additional effect resulting in a decrease in gain beyond trapping has been reported in 300 µm LGADs and is being investigated on 50 µm LGAD.
Investigate restoring gain loss after irradiation with increasing the bias voltage.
Pursue program to improve radiation tolerance by replacing Boron with Gallium.
Effect of trapping:
gain holes disappear,
looks like a change in gain

## Slide 16
Sentaurus Simulations (Julie Segal, SLAC)
LGAD #1 (low gain) 50 m thick Current vs. Time for 1 MIP @ 100 V Bias No radiation, 1e14, 5e14
LGAD #3 (high gain) 50 m thick Current vs. Time for 1 MIP @ 100 V Bias no radiation, 1e14, 5e14
N.B. “1 MIP” relates to the amount of charge is injected into the backside of the LGAD: maximum .
We are now measuring LGAD which were irradiated with protons at Los Alamos and neutrons at Llubljana.
Similar results are reported by many groups: TCAD works “right out of the box”

## Slide 17
HGTD Simulation, Design and fabrication.
G. Pellegrini, M. Carulla, S. Hidalgo, D. Flores, D. Quirion

## Slide 18
Design of HGTD ATLAS LGAD
3x3 mm2 and 2x2mm2 pads separated by 33 m (optimized by simulation).
P-stop isolation.
50 m thick active area. Support wafer can be etched away leaving a mechanical structure for mounting the read out chip and wire bonding.
Each chip reads out 4 channels.
Gain is set by the thickness of the substrate (max bias expected 200 V due to the planar junction breakdown), a complete (technological and electrical) simulation has been performed using Sentaurus and Silvaco to find the parameter necessary for the fabrication. Expected gain 6-10
Fabrication steps are the same than standards LGAD (with the thermal steps and implant energy and dose to be optimized by simulation).
First fabrication in 4” SOI wafers. Fabrication started in February 2016, wafers should be ready by May 2016.
Then fabrication should be optimized in 6” wafers. At the moment CNM is already doing a run of LGAD detectors in 6” wafers, 200 m thick. The run is due in 2016.
Open question: are they radiation hard? Fluences expected up to 5E15neq/cm2.

## Slide 19
Core Structure
Multiplication dopant implant as measured by SIMS. The fabrication process is the same as for the standard 300 m thick wafers.
50um
Electric Field @ 140V
50um
300um
3000um
30um
p-stop
Al
p+
n+
p-
p++
FE
passivation
SiO2
Al
Wet etching of silicon

## Slide 20
I-V Characteristics of thin detectors
50um epi
75um epi
Run 6474
300um FZ
20
Beware of the dimensions of the active volume

## Slide 21
Transient Simulations and Measurements

## Slide 22
CNM Wafer Layout
A1=8x8 matrix, 3mm pad
B1=8x8 matrix, 2mm pad
A2=4x4 matrix, 3mm pad
B2=4x4 matrix, 2mm pad
A3=2x2 matrix, 3mm pad
B3=2x2 matrix, 2mm pad
A4=2x2 matrix, 3mm pad, PIN
B4=2x2 matrix, 2mm pad, PIN
C=CT-PPS
D=TOTEM detector
Different diodes and test structures.
3x3 mm2 pixels.
Hole in the metallization
JTE only in the guard ring.
P-stop isolation.

## Slide 23
A. Anker, J. Chen, V. Fadeyev, P. Freeman, Z. Galloway, B. Gruey, H. Grabas, L.Hibbard, C. Labitan, Z. Liang, R. Losakul, S. N. Mak, C. W. Ng,
Hartmut F.-W. Sadrozinski, A. Seiden, N. Woods, A. Zatserklyaniy SCIPP, Univ. of California Santa Cruz, CA 95064, USA
B. Baldassarri, N. Cartiglia, F. Cenna, M. Ferrero
Univ. of Torino and INFN, Torino, Italy
G. Pellegrini, S. Hidalgo, M. Baselga, M. Carulla,
P. Fernandez-Martinez, D. Flores, A. Merlos, D. Quirion
Centro Nacional de Microelectrónica (CNM-CSIC), Barcelona, Spain
Contributors

## Slide 24
Conclusions
We took data in a triggered  source with 75 m thick epi LGAD with gain G ≈ 10 and measured a jitter of less than 40 ps.
The timing resolution in beam and laser tests are confirmed by the Weightfield 2 simulations.
We are developing specifications for LGAD to be used in the ATLAS HGTD.
We are looking forward to the next production run of UFSD at CNM available in May 2016.
Thanks to the organizers for arranging this year’s “Trento” meeting in Paris!