## Slide 1
Progress report on the analysis of PICOSECOND-MICROMEGAS test beam and calibration data
techniques and studies
Precise timing workshop
RD51 Mini-Week
21-02-2017
S. E. Tzamarias
Nuclear & Particle Physics Lab
Aristotle University of Thessaloniki

## Slide 2
Reminder: Arrival Time Estimation with test beam data
Reminder: SPE Charge Distribution
Comparison of SPE pulses with the muon data
Interpretation of the muon charge distribution in terms of SPE’s
Use the tracking information for the charge distributions
Misalignment, a naïve MC and a nice fit
The laser calibration: some observations
slewing and resolution vs e-peak amplitude
the resolution of the unobserved events
time resolution as jitter of the rise time

## Slide 3
REMINDER: Waveform Analysis
Raw signal
Guard Band
Ts
Te
Tp
Use the Fourier Filtered (F.F.) Waveform to define Ts, Tp, Te
Peak-position=Tp-Ts, Peak-asymmetry=(Tp-Ts)/(Te-Tp)
Several estimations of e-Peak Charge
using the raw or the F.F. signal
integrate the waveform from Ts to Tp or from Ts to Te
Integrate the whole waveform, raw or F.F., to estimate the total charge
Fourier Filtered signal

## Slide 4
Four +1 Fitting Strategies give the same results Timing at C.F. (20% of the Peak)
Global Sigmoid
Earlier Sigmoid
(the inflection point is also a middle point)
Polynomial Fit up to the inflection Point
Linear Interpolation
An extra strategy… Fit based on the average shape
See later….
REMINDER: FITS

## Slide 5
Measure and parameterize Slewing and Resolution as functions of the e-peak amplitude
The Pull Distribution is Normal with mean and sigma consistent with 0 and 1 respectively
ΔTcorrected/σ(Vpeak)
Almost a 1/sqrt(N) behavior
REMINDER: ~44 ps per muon

## Slide 6
Reject calibration events with >1 non-synchronous photons
Fit the SPE charge distribution with Polya (Gamma) Distribution
Analyze the “flame” calibration data

## Slide 7
After rejecting calibration events with >1 non-synchronous photons, the charge distribution is consistent with an SPE spectrum
REMINDER:

## Slide 8
Calibration and Comparison
at RUN284 Operating Parameters
(Vmesh =+450 and Vdrift=-350)
SPE – Calibration RUN

## Slide 9
The e-Peak shape in SPE-Calibration and Muon Data

## Slide 10
There is NO significant Background Contribution to the Data

## Slide 11
In comparison with SPERMS~0.9
In comparison with μRMS~ 7.3
There is not any significant broadening of the e-Peak Charge Distributions

## Slide 12
Different e-Peak Charge Estimations result to the same Distribution Parameters in both SPE-Calibration and Muon Data

## Slide 13
Different e-Peak Charge Estimations result to the same Distribution Parameters in both SPE-Calibration and Muon Data

## Slide 14
How can we estimate the mean number of PEs produced by the Muons of the Beam
How we interpret the spread of the e-Peak Charge Distribution in terms of the Detector Response to a Single PE

## Slide 15
SPE: mean charge and variance
PE Multiplicity: mean and variance
Mean and variance of the muon charge in terms of the above

## Slide 16
Example: using F.F. e-peak charge
and the Polya Fit parameters
we estimate depending of the μ parameters used
According to this analysis
the data REJECT the case of a POISSONIAN PE distribution
E[Q] and V[Q] from the muon data whilst Qe and Ve from the Polya Fit to SPE calibration data
We estimate the mean number of observed PEs and the variance of their multiplicity distribution

## Slide 17
The curve is the result of a fit
Convolution of a Spe- Polya (with parameters defined by the SPE calibration) with a Poissonian for the SPE multiplicity
The mean of the Poissonian is the only free parameter

## Slide 18
Two Components Fit
major component: 60%
Convolution of a Spe- Polya (with parameters defined by the SPE calibration) with TWO Poissonians for the SPE multiplicity
The mean of the Poissoniana are only free parameter

## Slide 19
Number of Tracks
Beam Spatial Distribution wrt the 5x5 mm2 scintillator

## Slide 20
LOW
High
Mean e-Peak Charge per Track
X Track-Impact Coordinate
Y Track-Impact Coordinate
Weighted Beam Spatial Distribution: weight each track by the corresponding e-peak charge
Mean e-Peak Charge Distribution: divide the “Weighted Beam Spatial Distribution” by the “Beam Spatial Distribution’
Calculate the statistical errors

## Slide 21
LOW
High

## Slide 22
“Low Beam” Region
“High Beam” Region
The timing resolution is very different in the two regions

## Slide 23
LOW
High
Mean No of PEs per Track
X Track-Impact Coordinate
Y Track-Impact Coordinate
Divide by the mean spe Charge determined by the Polya Fits

## Slide 24
Simulation Studies
simulate the beam by “bootstraping” the RUN284 reconstructed tracks (plus a 0.2 mm jitter)
assume μCherenkov =12
perfect alignment, perfect efficiency (except “pillars losses”)
check “ high” and “low” regions
in the demonstration plot the pillars with 0.2 mm radius

## Slide 25
Accepted Pes Spatial Distributions:
all, b) pes from “low” beam region, c) pes from “high” beam region
the arcs represent the detector geometrical limits when displaced by (3.5, -3.5) and (4, -4) mm

## Slide 26
Does a possible misalignment explains the observed spatial dependence of the efficiency ?
pes from “low” beam region, pes from “high” beam region

## Slide 27
Simulation is performed with 0.1mm pillars radius, 2.5 mm from the anode periphery
Simulated Pattern
Data
Data Bins with less than 10 tracks have been put to zero and they do not participate in the fit
Mean No of PEs per Track
X Track-Impact Coordinate
Y Track-Impact Coordinate
i and j correspond to X and Y track impact coordinates
Dij are the Data bins, whilst eij are the corresponding statistical errors.
Sij are the simulated pattern bins which are functions of the mean number of pes produced (μCh) and the anode’s displacement (δx , δy).
Sij are produced by different simulation runs, selecting the parameters (δx , δy, μCh ) accordingly.
The estimation is performed by minimizing the following χ2 estimator, using two different minimization algorithms
where r(δx,δy) is a scale factor that can be a treated as a function of the
displacement or as a constant
FIT STRATEGY

## Slide 28
χ2 MINIMIZATION METHODS
Relative Scaling
Simulated Patterns are generated for several displacement points (δx , δy) using a preselected μCh , the same for all the patterns. The “true” displacements are estimated as the values that correspond to the simulated pattern with the minimum “reduced χ2”
The relative scaling corresponds to scaling the number of produced pes, independently, for each simulated pattern. Then, the estimated value of μCh is given by
Iterative Fit
Simulated Patterns are generated for several displacement points (δx , δy) using a preselected μCh , the same for all the patterns. The “true” displacements are estimated as the values that correspond to the simulated pattern with the minimum “χ2” :
A scale factor is calculated using the simulated pattern produced at :
If the scale factor is not 1 then update μ’Ch=rglobal μCh, simulate new patterns with μ’Ch and repeat the fit
When rglobal=1, the procedure have converged

## Slide 29
“Relative Scaling” FIT - RESULTS II
Full Scan
68.3% Level (minimum+2.3)
Minimum at δx=2.5mm , δy=-2.3mm
Approximate 68.3% limits 2.3mm<δx<2.6mm and -2.4mm<δy<-2.0mm
χ2
χ2

## Slide 30
-2 mm < Yb < -1 mm
-1 mm < Yb < 0 mm
0 mm < Yb < 1 mm
1 mm < Yb < 2 mm
2 mm < Yb < 3 mm
Iterative FIT - RESULTS III
Full Scan
The curve is the convolution of a SPE-Polya with the pe multiplicity distribution evaluated by the MC fit

## Slide 31
Systematical Errors:
Beam
PEs
Anode displaced relative to the beam at δx=2.55mm, δy=-2.1mm
The underlying model
Change the pillars orientation
Change the pillars radial coordinate
F=1.7
F=1.7
F=1.7

## Slide 32
Systematical Errors:
Anode displaced relative to the beam at δx=2.5mm, δy=-2.1mm
F=2.2
Anode inefficiency
( 0.5mm-wide outer ring)
Fit using a “reduced anode” model: 4.5 mm anode radius
Relative Scaling: χ2min=58.5 δx=2.1 δy=-1.8
Iterative Fit: χ2min=61 δx=2.1 δy=-1.7
It finds a different minimum with larger Chi2
But …
Compare with “full anode” results
χ2min=54 , 2.3mm<δx<2.6mm and -2.4mm<δy<-2.0mm @ 68.3% CL

## Slide 33
F=1.8
In good agreement with the spe results
The fit is much more sensitive to the “number of pes per track distribution” than to the real misplacement of the anode wrt the beam (i.e. we estimate an effective position !!!)
Fit Results using the “reduced anode” model
Very similar with the “full anode model”

## Slide 34
Calibration @ Other Working Points - PRELIMINARY
 |  |  | Muon Run |  | SPE Run – Polya Fit |  | Parameters | 
Run Number | PS2 Mesh | PS2 Drift | E[Q] | V[Q] | Qe | Ve | <Νpe-obs.> ± 0.5 | Θ ± 0.3 (??)
277 | +450 | -325 |  |  |  |  |  | 
278 | +450 | -300 |  |  |  |  |  | 
279 | +425 | -325 |  |  |  |  |  | 
280 | +425 | -300 |  |  |  |  |  | 
281 | +425 | -350 |  |  |  |  |  | 
282 | +450 | -275 |  |  |  |  |  | 
284 | +450 | -350 | 13.8 | 53.9 | 1.5 | 0.81 | 9.3 | 1.8
285 | +475 | -250 | 1.94± 0.02 Syst. error 1.75± 0.02 | 1.07 | 0.35± 0.04 | 0.04 | 5. | 2.2
286 | +475 | -275 | 3,4 | 3.24 | 0.51 | 0.1 | 6.7 | 1.6
287 | +475 | -300 | 6.51 | 12.5 | 0.8 | 0.29 | 8.1 | 1.2
288 | +475 | -325 | 13.5 | 53.7 | 1.36 | 0.81 | 9.9 | 1.3
289 | +450 | -325 |  |  |  |  |  | 

## Slide 35
Studies with the Laser – Calibration Data

## Slide 36
Very good agreement between the different timing methods

## Slide 37
?
Single Gaussian Fit
RUN 25-01-2017_CF4-0.2C2H6 600-500
0.04V < Ve-peak < 0.05V
0.25V < Ve-peak < 0.30V

## Slide 39
0.04V < Ve-peak < 0.05V
0.25V < Ve-peak < 0.30V

## Slide 40
600-425

## Slide 41
CF4/600 V
Time Resolution dependence on the e-peak amplitude
600-425
600-450
600-475
600-500

## Slide 42
Corrected for trigger offsets
Slewing dependence on the e-peak amplitude

## Slide 43
Global Corrections
All the RUNs-600 treated using the “global” slewing and resolution parameterizations

## Slide 44
600-425
We should correct the resolution curves for threshold effects

## Slide 45
Data collected with different thresholds
Predicted resolutions for a global 0.018 V pulse acceptance threshold
How we predict?

## Slide 49
 | 600-425 Cut: 0.018 V | 600-450 Cut: 0.018 V | 600-475 Cut: 0.04 V | 600-500 Cut: 0.05 V
data | 107±1 ps | 94.5±0.7 ps | 70.0± 0.5 | 62.2±0.6
“prediction” | 102 ps | 96.2 ps | 70.6 ps | 61.4 ps
The errors are from gaussian fits after slewing correction
The “prediction” is the convolution of the Polya parameterization of the PH distribution with the resolution vs PH parameterization

## Slide 50
In the following Rise Time =1/R3

## Slide 51
RUN 25-01-2017_cf4-0.2c2h6_600-500
0.05V < Vp < 0.06V
0.10V < Vp < 0.11V
0.18V < Vp < 0.20V
0.25V < Vp < 0.35V

## Slide 53
Remember that

## Slide 55
Comparison of the “Pull” Distributions (with and without slewing corrections) to Standard Gaussians

## Slide 56
REMINDER: Estimate the e-Peak Charge Distribution Parameters
Parameterize the PicoMM response to N Pes by Gamma Distribution Functions

## Slide 57
SPE Calibration Data
Muon Data
We have developed criteria to reject multi-photon calibration events
REMINDER:

## Slide 58
Let us assume that the single pe’s charge distribution is given by:
where
The pdf for the charge which is produced by n pe’s (n>0) will be:
Because , where follow (1) and are mutually independent, then
(1)
(2)
(3)
(4)
Furthermore, the pedestal distribution which corresponds to 0-pes, has

## Slide 59
Consider now muons producing n pes, where the variable n follows some pdf , where
(5)
and is a vector of other parameters (this distribution could be Poissonian)
Then the total accumulated charge (that is proportional to the integral of the corresponding waveform) is distributed according to:
(6)
The expectation value of the total charge is given by
(7)

## Slide 60
(8)