## Slide 1: Completion of DYB near hall RPC gas system installation
Changguo Lu, Princeton University
Sun Hansheng, Guan Mengyun, Xu Jilei, IHEP
Viktor Pec, Charles University

## Slide 2: Schedule
Our original gas system installation plan from July 14, 2010 Muon installation workshop:
Done, November 5-th, 2010, as scheduled
Done, March 31, 2011, as scheduled.
Remaining things:
Replace exhausting gas tubing with thicker pipe,
Debugging/Monitoring digital bubbler readout and gas chromatograph system.

## Slide 3: What have been done for RPC gas system installation since last November
Installed gas distribution crates on RPC module supporting frame;
Connected gas tubing from modules to distribution crates;
Finished up the cable trays for gas room gas pipe, cables, power cores, etc.
Connected fire water pipe to the gas cabinet and tested;
Posted signs on Isobutane gas room door, gas cabinet, fire water pipes, valves, locked fire water valves to meet safety requirement;
Established the communication between host pc in the gas room and digital bubbler readout crate on the RPC module supporting frame;
Calibrated all four gases with Daya Bay near hall GC. Repaired Linao near hall GC (left over from previous vender’s installation) and found new defect part for this GC;
DGUT group has installed the detector control system interface box.

## Slide 4: What have been done for gas system installation(cont’d)
6 gas system procedures and 5 checklists have passed ORR review and translated into Chinese and conducted 2 gas system training sessions on site;
Got approval from LSO to bring Isobutane cylinder into gas room, successfully mixed Daya Bay RPC gas mixture and flow through all 108 channels;
Tested gas system in full flow range (2 volumes/day). Found the original exhausting gas tubing too small that needs to be replaced.

## Slide 5: Gas system procedures and checklists
DocDB #6071 - #6076:
DocDB #6071 Procedure for replacement of the Isobutane cylinder w/ checklist;
DocDB #6072 Procedure for replacement of the SF6/R134A cylinder w/ checklist;
DocDB #6073 Procedure for replacement of the argon cylinder w/ checklist;
DocDB #6074 Procedure for the inspection of the RPC gas system safety device w/ 2 checklistss;
DocDB #6075 Procedure for the gas system routine leak check;
DocDB #6076 Procedure for operation of the RPC gas system.

## Slide 6: Cable tray in gas room
Was
Now
Cable tray installed

## Slide 7: Cable tray (cont’d)
Was
Now

## Slide 8: GC calibration chromatograms
Ar
SF6

## Slide 9: GC calibration chromatograms (cont’d)
R134A
Isobutane

## Slide 10: Daya Bay RPC gas mixture
Raw data. Corrected by calibration curves, the mixing ratio is:
65.73: 3.95: 29.81: 0.51
Very close to our nominal ratio: 65.5:4:30:0.5.

## Slide 11: Gas distribution/digital bubblers
Was temporarily mounted on a rack in the gas room:
Now they are mounted on the supporting frames:

## Slide 12: Digital bubbler readout
The screen capture of the digital bubbler readout histograms for all 7 crates (16 channels/crate):
Three screen captures show the RPC “breathing” in the gas flow. If we connect the bubblers w/o RPC chambers we’ll see the bubbling rate histograms w/o “breathing” as shown in next slide.

## Slide 13: Bubbler histograms w/o “breathing”
Connect the outlet of the gas distribution channels to the digital bubblers directly (w/o RPC chambers), the bubbling rate histograms are quite uniform, no “breathing”.

## Slide 14: Gas switchover panels, manifolds, pipes & cabinet
Water pipe to drain the fire water once every three months.
Fire water pipe valve has been locked on open position.

## Slide 15: Remaining jobs for the first Daya Bay RPC gas system installation
Thoroughly check the leakage of the gas system once more;
Further improving the exhausting gas tubing with 1~2” thick pipe;
Work with DGUT group/DCS group to display gas system status on DCS screen;
Running bubbler readout GUI program, check its stability;
Running GC sequence on a separate PC to check its stability;
Provide Daya Bay RPC gas mixture to RPC modules and commissioning the RPC system.

## Slide 16: Special thanks to Strong support from LIM Xiaonan; US/China LSO/SO; IHEP gas installation team; Other collaboration members, who participated in the gas system installation.